# Converter Installation and Wiring Diagram

- VIR-832  Installation Instruction

- Wall mounting: Fix the converter with 4 Φ6 expansion bolts or normal nails.

- DIN-rail mounting by using rail fixing clamps.

# VIR-832 Wiring Diagram

## VIR-832(M)-24V Module type

# Transducer Introduction and Wiring Diagram

## Clamp on type transducer

# VIR-832 Converter Installation and Transducer Wiring Guide

## Converter Installation

- For wall mounting, secure the VIR-832 converter using four Φ6 expansion bolts or standard nails.

- For DIN-rail mounting, use rail fixing clamps to attach the converter.

## VIR-832 Wiring

- VIR-832(M)-24V is a module type and includes specific wiring considerations as per the diagram.

## Transducer Types and Wiring

- Clamp on type transducer: Includes an introduction and wiring diagram for proper installation.

- Insertion type transducer (5.2): Comes with an introduction and accompanying wiring diagram.

- Inline type transducer (5.3): Features an introduction and wiring diagram for the inline model.

## Display and Operation

- The display consists of a 2×20 character LCD with backlight.

- Settings are available for adjusting backlight time and contrast.

- A keyboard is provided for operation and configuration.

- Wiring Diagram

- Insertion transducer

- Wiring Diagram

## Inline type transducer:

- Wiring Diagram

# Display and Operation

## Display is 2×20 characters LCD with backlight, available to set backlight time and contrast.

- 16-key Keyboard-VIR-832

0  -  9 and  . are used for inputting numbers or menu numbers.

◄ is used for back left or delete the left character.

▲/+ and ▼/- are used for entering into the last and next menu. Also can be used as ± sign when inputting numbers.

MENUis used for accessing the menu. Press this key first, then type the number keys to enter into the matched menu.

ENTis the ENTER key, used for confirming the contents you input or choose.

## Separated Mounting

- 4-key Keyboard	-VIR-832-M

## VIR-832-M

## MENU : used for entering into menus.

## ▲: used for menuup or choosing 0-9, +, -

## : used for menudown or moving the cursor to next.

## ENT : used for finishing menu inputting or entering into submenu.

## Operation:

## The user interface of this flow meter comprises about 100 different menu windows that are numbered by M00, M01, M02 … M99.

## Method to enter Menu: Press MENU first, and follow the two-digit number keys. Take M35 as an example, the correct key sequence is MENU35

## To move between the adjacent menus, press ▲/+ and ▼/- for 16-key keyboard; press ▲and for 4-key keyboard.

## Menu Details

## Quick setup of measured parameters

## Accurate measured parameters can have a great influence on measuring precision and reliability. It is suggested to measure the practical perimeter and wall thickness of the pipeline. Ultrasonic thickness gauge can be used to measure the pipe thickness.

## Measured parameters setup is from Menu10 to Menu29. Please complete one by one.

## Following parameters need to be inputted before measurement:

- Outer diameter  unit: mm

- Pipe thickness  unit: mm

- Pipe material

- Lining parameters: thickness and sound velocity (If have lining)

- Liquid type

- transducer type

- transducer mounting type

## Above parameters setup generally follow the steps below:

## 1.Press keys MENU 1  1  to enter M11 window to input the pipe outer diameter, and then press

## ENT key.

## 2.Press key ▼/- to enter M12 window to input the pipe outer diameter and then press ENT key.

## 3.Press key ▼/- to enter M14 window, and press ENT key to enter the option selection mode.

## Use keys ▲/+ and ▼/- to select the pipe material, and then press ENT key.

## 4.Press key  ▼/- to enter M16 window, press ENT key to enter the option selection mode. Use

## keys ▲/+ and ▼/- to select the liner material, and then press ENT key. Select “No Liner”,

## if there is no liner.

## 5.Press key ▼/- to enter M20 window, press ENT key to enter the option selection mode. Use

## keys ▲/+ and ▼/- to select the proper liquid, and then press ENT key.

## 6.Press key ▼/- to enter M23 window, press ENT key to enter the option selection mode. Use

## keys ▲/+ and ▼/- to select the proper transducer type, and then press ENT key.

## 7.Press key ▼/- to enter M24 window, press ENT key to enter the option selection mode. Use

## keys ▲/+ and▼/- to select the proper transducer mounting method, and then press ENT key.

## 8.Press key ▼/- to enter M25 window and get the transducer installation distance.

## 9.Press MENU 2  6  to store the parameters setup.

## Choose installation points

## Proper installation point is a key for transducer installation. Following factors must be considered: Full filled pipeline, shaking, steady flow, scaling, temperature, pressure, EMI, instrument well.

## Full filled pipeline

## Following situations can be full filled of liquid:

## Vertical upward               Obliquely upward                  Lowest point

## > Vibration

## There cannot be obvious vibration on the installation point, otherwise it needs to be tightened.

## >>Steady flow

## Steady flow is helpful for ensuring measurement accuracy.

## Standard requests for steady flow are:

## 1. The pipe should be far away from pump outlet and half-open valve.

## 10D to upstream and 5D to downstream. (D means outer diameter)

## 2. 30D to pump outlet and half-open valve.

## Scaling

## The inside scaling would have bad effect on ultrasonic signal transmission, and would   decrease the inner diameter as well. As a result, the measurement accuracy can not be guaranteed. Please try to avoid choosing the installation point with inside scaling.

## Temperature

## The liquid temperature on installation point should be in the working range of transducers. Please try to choose the point with lower temperature. Avoid to choose points like the outlet of boiler water and heat exchanger. Return water pipe would be better.

## Temperature range of standard clamp on and insertion transducers: -30 ~ 90℃

## Temperature range of high temperature clamp on and insertion transducers: -30 ~ 160℃

## Pressure

## The maximum pressure for standard insertion and inline transducer is 1.6MPa

## Out of this range need customized.

## 7.2 Clamp on transducer Installation

## Before installation, please verify the parameters of pipeline and liquid. To ensure the installation accuracy.

- Installation procedure

## Select an installation method → Input the measuring parameters → Clean pipe surface →

## Install transducers → Check the installation

- Select an installation method

## There are two different methods for clamp on transducers: V method and Z method.

## >> V method

## V method should be priority selected for pipe sizes DN25 - DN200. Let the pair of transducers horizontal alignment, the central line in parallel with the pipeline axis.

## >> Z method

## Z method should be priority selected for pipe sizes DN200 - DN6000. Also can be used when V method doesn’t work well. Make sure the vertical distance of two transducers equals to the installation distance, and the two transducers are on the same axis surface.

- Positioning installation points

## >> V method

## The line between two transducers is parallel to pipe axis, and equal to the distance shown in the converter. As shown, A, B are the two installation points.

## >> Z method

## Firstly according to the installation distance shown in converter, positioning two points A, C

## on the same side of pipeline. AC is parallel to pipe axis.

## Perpendicular to the pipe axis, opposite to point C, get Point B.

## Check. Measure the length between A and B from both sides of the pipe, get AB1 and AB2. If

## AB1 = AB2, then B is the correct point. If not, need to positioning point B and C again.

## As shown, A, B are the two installation points.

- Clean the surface of installation points

## Paint, rust and anti-corrosive coating on installation points need to be cleaned. It’s good to use a polishing machine to get the metal luster.

## As shown below:

- Install transducers

## After transducer wiring and sealing, please evenly smear 2-3mm couplant on the transducer emitting surface. Then put the transducers on the installation points, fixed with steel belt or steel rope.

- Check Installation

## Please see details in Chapter 7.5

## 7.3 Insertion type transducer installation

## Before installation, please verify the parameters of pipeline and liquid. To ensure the installation accuracy.

- Installation procedure

## Select an installation method → Input the measuring parameters → Positioning installation points → Fix ball valve base →Open hole under pressure →Install transducers → Check the installation

- Select installation method and positioning installation points

## Insertion type transducers are suitable for pipe sizes > 50mm.

## Two different installation methods: V method and Z method. Generally use Z method, only use V method for lack of space.

## >> V method

## V method can be used for DN50mm - 300mm. Let the pair of transducers horizontal alignment, the central line in parallel with the pipeline axis, and the transmit direction mush be opposite.

## >> Z method

## Z method can be used for all pipes > DN50mm. Make sure the vertical distance of two transducers equals to the installation distance, and the two transducers are on the same axis surface. The transmit direction mush be opposite.

## >> Parallel insertion

## If there is insufficient installation space or the transducers can be only installed on the top of pipeline, parallel insertion transducer will be a good choice. (Pipe size ≥300)

## Positioning of parallel insertion transducer need to meet the 3 factors as follow:

- Installation distance = Vertical distance of two transducers along the pipe axis direction

- Make sure two transducers are in the same horizontal line,

## Insertion depth = 1/3 inner diameter

- Users can set the distance between transducers by themselves. Recommend 300~500mm

- Positioning installation points

## >> V method

## The line between two transducers is parallel to pipe axis, and equal to the distance shown in the converter. As shown, A, B are the two installation points.

## >> Z method

## Firstly according to the installation distance shown in converter, positioning two points A, C

## on the same side of pipeline. AC is parallel to pipe axis.

## Perpendicular to the pipe axis, opposite to point C, get Point B.

## Check. Measure the length between A and B from both sides of the pipe, get AB1 and AB2. If

## AB1 = AB2, then B is the correct point. If not, need to positioning point B and C again.

## As shown, A, B are the two installation points.

- Fix ball valve base

## >> Welding Fix

## For carbon steel pipes, the ball valve base can be welded directly. Make sure that the central point of ball valve base is overlapped with the transducer installation point.

## Matters need attention:

- Please take the PTFE sealing gasket out from the base before welding.

- Please clean the pipe surface around welding point before welding. Pay attention that there should not be any air hole during welding, which can avoid leaking. Welding strength must be ensured.

- Do not sputter welding slag on the base thread.

- Non-deformation of base during welding.

## After welding, tighten ball valve into the base.

## >> Pipe hoop Fix

## For pipes can’t be welded directly like cast iron pipe, cement pipe, copper pipe and composite pipe, customized pipe hoop is recommended.

## The hoop center should be overlapped with the transducer installation point. Please compress the sealing gasket tightly to avoid leaking.

- Open hole

## After finishing the installation of ball valve and base, insert the open-hole tool into ball valve and lock it. Then open the ball valve, start drilling, from slow to fast. Close ball valve after drilling.

## See more details in the video of insertion transducer installation.

- Install transducer and adjustment

## Adjust the proper insertion depth and transmit direction to get good ultrasound signal.

## >> Insertion depth adjustment

## Adjust the depth scale according to pipe wall thickness, and completely push in the transducer rod.

## >> Transmit direction

## There is a indicating arrow on the transducer junction box, the arrow direction on two transducers should be opposite and parallel to the pipe axis.

## >>Operation steps

- Tighten the locknut into ball valve, adjust the insertion depth scale.

- Open ball valve, completely push in the upstream transducer rod. Adjust the transmit direction parallel with pipe axis, and point to the installation point of downstream transducer. Lock it after adjustment.

- Install downstream transducer in the same way. Adjust the transmit direction to get the best signal strength and watching Menu91, if the value is between 97% ~ 103%, the installation is correct. If not, need to re-adjust the insertion depth and transmit direction until meet the requirement.

- Check installation

## Please see details in Chapter 7.5

# 7.4 In-line type transducer installation

## After choosing the installation point, install the transducer in the pipeline with companion flanges. Then connect the transducer to converter with special signal cable. Installation is complete.

- Installation method

- Check installation

## Please see details in Chapter 7.5

## 7.5 Check Installation

## The flow meter includes the detection ability. M90 is used for checking signal strength and quality. M91 is used for checking the ratio of measured and theoretical transmission time (transmission time ratio).

- Check signal strength and quality

## M90 is used for checking the signal strength and signal quality(Q value) of upstream and downstream transducers.

## Signal strength is represented by numbers 00.0 ~ 99.9, 00.0 means no signal and 99.0 means maximum signal. Generally, the flow meter can work properly when signal strength is > 60.0

## Signal quality (Q value) is represented by numbers 00 ~ 99. 00 means signal is worst and 99 means signal is best. The flow meter can work properly when Q > 60.

## During the installation, please adjust the transducer to make the signal strength and signal quality the larger the better. This will ensure the flow meter long term stable operation and lead to accurate measurement.

| ## Signal strength and Q value | ## Installation Judgement |
| --- | --- |
| ## < 60 | ## Can not work |
| ## 60~75 | ## Bad |
| ## 75~80 | ## Good |
| ## >80 | ## Excellent |

- Check transmission time ratio

## M91 is used for displaying transmission time ratio. It is a percentage ratio between theoretical transmission time and measured transmission time. It shows the relation between setting parameters and actual transducer installation distance. This ratio should be between 97% ~ 103%.

## If not in the range of 97%~103%, it means that the parameters and transducer installation distance are inconsistent. Please check separately.

- Finish Installation

- Commonly used menus. M00 or M02 is for meter reading. M30~M33 is for unit selection. M40 is for selecting damping factor, generally 5~10 sec. M60 is for correcting time and date. M26 is for curing parameters.

- To avoid signal reduction and improve anti-jamming ability, it is better to use the customized signal cable from flow meter manufacturer.

- The length of cables between converter and transducer should be as short as possible, cannot exceed 200m.

## 4) The temperature and humidity of working environment should be in the range of technical specifications. Avoid direct sunlight on LCD.