## Sheet1
Menu No. | Function
M00 | Display flow rate and NET totalizer.
If the net totalizer is turned off(refer to M34), the net totalizer value shown on the screen is the value prior to its turn off.
Select all totalizer unit in menu M31.
M01 | Display flow rate, velocity.
M02 | Display flow rate and POS(positive) totalizer.
If the positive totalizer is turned off, the positive totalizer value shown on the screen is the value prior to its turn off.
M03 | Display flow rate and NEG(negative) totalizer.
If the negative totalizer is turned off, the negative totalizer value shown on the screen is the value prior to its turn off.
M04 | Display date and time, flow rate. The date and time setting method is found in MENU60.
M05 | Display energy rate(instantaneous Caloric)and total energy (Caloric).
M06 | Display temperatures, inlet T1, outlet T2.
M07 | Display analog inputs, AI3/AI4, current value and its corresponding temperature or pressure or liquid level value.
M08 | Display all the detailed error codes.
Display working condition and system error codes. ‘R’ stands for normal; others refer to Chapter 5 for details.
M09 | Display today’s total NET flow.
M10 | Window for entering the outer perimeter of the pipe.
If pipe outer diameter is known, skip this menu and go to Menu 11to enter the outer diameter.
M11 | Window for entering the outer diameter of the pipe. Valid range:0 to 18000mm.
Note: you just need to enter either the outer diameter in M11 or the perimeter in M10.
M12 | Window for entering pipe wall thickness
You may skip the menu and enter inner diameter in M13 instead.
M13 | Window for entering the inner diameter of the pipe
If pipe outer diameter and wall thickness are enter correctly, the inner diameter will be calculated automatically, thus no need to change anything in the window
M14 | Window for selecting pipe material
Standard pipe materials (No need to enter material sound speed ) include:
(0) carbon steel (1) stainless steel (2) cast iron (3) ductile iron (4) copper (5) PVC (6) aluminum ,(7) asbestos (8) fiberglass
(9) other( need to enter material sound speed in M15)
M15 | Window for entering the pipe material speed, only for non-standard pipe materials
M16 | Window for selecting the liner material, select none for pipes without any liner.
Standard liner materials(no need to enter the liner sound speed) include:
(0) None, No liner (1) Tar Epoxy (2) Rubber (3) Mortar (4) Polypropylene (5) Polystryol (6)Polystyrene (7) Polyester (8) Polyethylene (9) Ebonite (10) Teflon
(11) Other (need to enter liner sound speed in M17)
M17 | Window for entering the non-standard liner material speed.
M18 | Window for entering the liner thickness, if there is a liner
M19 | Window for entering the ABS thickness of the inside wall of the pipe
M20 | Window for selecting fluid type
For standard liquids(no need to enter fluid sound speed) include:
(0) Water (1) Sea Water (2) Kerosene (3) Gasoline (4) Fuel oil (5) Crude Oil
(6) Propane at -45C (7) Butane at 0C
(8)Other liquids(need to enter sound speed in M21 and viscosity in M22)
(9) Diesel Oil (10)Caster Oil (11)Peanut Oil (12) #90 Gasoline (13) #93 Gasoline (14) Alcohol (15) Hot water at 125C
M21 | Window for entering the sound speed of non- standard liquid, used only when option item 8 ‘Other’ is selected in M20
M22 | Window for entering the viscosity of the non-standard liquids, used only when option item 8 ‘Other’ is selected in M20
M23 | Window for selecting transducer type, There are 22 types as following
0. Standard-M (The middle size) 1. Insertion Type C
2. Standard-S 3. User Type 4. Standard B
5. Insertion Type B(45) 6. Standrad-L (The large size transducers)
7. JH-Polysonics 8. Standard-HS (small size transducer for TUF-2000H
9. Standard-HM (middle size transducer for Handheld flow meter)
10. Standard-M1 (middle size transducer #1)
11. Standard-S1 (small size transducer #1)
12. Standard-L1 (large size transducer #1)
13. PI-Type 14. FS410 (middle size transducer for FUJI flow meter)
15. FS510 (large size transducer for FUJI flow meter)
16. Clamp-on TM-1 (Middle size transducer for Taosonics Instrument)
17. Insertion TC-1 (for Taosonic Instrument)
18. Calmp-on TS-1 (small size for Taosonics Instrument)
19. Calmp-on TS-1 20. Clamp-on TL-1 (For Taosonics Instrument)
21. Insertion TLC-2 (For Taosonics Instrument)
22. Clamp-on M2 23. Clamp-on L2
If the user-type-transducer is selected, you need enter additional 4 user-type-wedge parameters that describe the user transducers.
If the PI-type transducer is selected, you need enter additional 4 PI-type transducer parameters that describe the PI-type transducers
M24 | Window for selecting the transducer mounting methods
Four methods can be selected:
(0) V-method (1) Z-method (2) N-method (3) W-method
M25 | Display the transducer mounting spacing or distance
M26 | (1) A switch for the parameters in flash memory will be loaded when power is turned on. The default option is that the parameters will be loaded. If this switch is not turned on, the system will try to use the parameters in the system RAM, if these parameters are ok, otherwise the system will load the parameters in flash memory
(2) Function to store the current parameters into the flash memory, so that these parameters will be solidified and will be loaded as the default parameters every time when power is turned on.
M27 | Entry to store to or restore from the internal Flash memory, as many as 9 different pipe parameter configurations
To save or load the current setup parameter, use the going up or going down keys to change the address number, press ‘ENT’ key, and use going down or going up keys to select to save to or load from the memory.
M28 | Entry to determine whether or not to hold (or to keep) the last good value when poor signal condition occurs. YES is the default setup.
M29 | Entry to setup empty signal threshold. When the signal is less than this threshold, the pipe is regarded as empty pipe, and the flow meter will not totalize flow.
This is based on the fact that, for most occasions, when pipe is empty, the transducer would still receive signal, just smaller than normal, As a result, The flow meter would show normal operation, which is not correct.
Make sure that the entered value must be less than the normal signal strength.
When much noisy signals are received, to make sure the flow meter will not incorrectly totalize flow, there is also a ‘Q’ threshold should be entered in M.5
M30 | Window for selecting unit system. The conversion English to Metric or vice versa will not affect the unit for totalizers.
M31 | Window for selecting flow rate unit system.
Flow rate can be in
0. Cubic meter short for (m3)
1. Liter (l)
2. USA gallon (gal)
3. Imperial Gallon (igl)
4. Million USA gallon (mgl)
5. Cubic feet (cf)
6. USA liquid barrel (bal)
7. Oil barrel (ob)
The flow unit in terms of time can be per day, per hour, per minute or per second. So there are 32 different flow rate units in total for selection.
M32 | Window for selecting the totalizers unit. Available units are the same as those in M31
M33 | Window for setting the totalizer multiplying factor
The multiplying factor ranges from 0.001 to 10000. Factory default is 1
M34 | Turn on or turn off the NET totalizer
M35 | Turn on or turn off the POS (positive) totalizer
M36 | Turn on or turn off the NEG(negative) totalizer
M37 | (1) Totalizer reset
(2) Restore the factory default settings parameters. Press the dot key followed by the backspace key. Attention, It is recommended to make note on the parameters before doing the restoration
M38 | Manual totalizer used for easier calibration. Press a key to start and press a key to stop the manual totalizer.
M39 | Language selection.
The selection could also be changed automatically by the system, if English LCD display is used as the display device.
M3A | Setup for local segmental LCD display. Enter 0 or 1 for the non-auto-scan mode; Enter 2~39 for the auto-scan mode. In the auto-scan mode the display will automatically scan displaying from 00 to the entered number of the local segmental LCD display.
M40 | Flow rate damper for a stable value. The damping parameter ranges from 0 to 999 seconds.
0 means there is no damping. Factory default is 10 seconds
M41 | Low flow rate (or zero flow rate) cut-off to avoid invalid accumulation.
M42 | Zero calibration/Zero point setup. Make sure the liquid in the pipe is not running while doing the setup.
M43 | Clear the zero point value, and restore the solidified zero point value.
M44 | Set up a flow bias. Generally this value should be 0.
M45 | Flow rate scale factor. The default value is ‘1’.
Keep this value as ‘1’, when no calibration has been made.
M46 | Networks address identification number. Any integer can be entered except 13(0DH, carriage return), 10 (0AH, line feeding), 42 (2AH), 38, 65535.
Every set of the instrument in a network environment should have a unique IDN. Please refer to the chapter for communication.
M47 | System locker to avoid modification of the system parameters.
If password is forgotten, you could send a command ‘LOCK0’ to the serial input to unlock. Or you can write 0 to REGISTER49-50 under MODBUS protocol.
M48 | Entry to linearity correcting data inputs. By using of this function, the non-linearity of flow meter will be corrected. Correcting data shall be obtained by careful calibration.
M49 | Displays the input contents for the serial port.
By checking the displays, you can know if the communication is ok.
M50 | Switches for the built-in data logger. There are as many as 22 different items can be chosen. To turn this function, select ‘YES’ the system will ask for selecting the items. There are 22 items available. Turn on all those items you want to output
M51 | Window to setup the time of scheduled output function (data logger, or Thermo-printer). This includes start time, time interval and how many times of output. When a number great than 8000 entered for the times of output, It means the output will be keeping always. The minimum time interval is 1 second and the maximum is 24 hours.
M52 | Data logging direction control.
(1) If ‘Send to RS485’ is selected, all the data produced by the data logger will be transmitted out through the RS-232/RS485 interface
(2) If ‘To the internal serial BUS is selected, the data will be transmitted to the internal serial bus which allows a thermal printer, or a 4-20mA analog output module, to be connected to it.
M53 | Display analog inputs, AI5, current value and its corresponding temperature or pressure or liquid level value.
M54 | Pulse width setup for the OCT (OCT1) output. Minimum is 6 mS, maximum is 1000 mS
M55 | Select analog output (4-20mA current loop, or CL) mode. Available options:
(0) 4-20mA output mode (setup the output range from 4-20mA)
(1) 0-20mA output mode (setup the output range from 0-20mA, This mode can only be used with Version-15 flow meter)
(2) RS232 Serial port controls 0-20mA
(3) 4-20mA corresponding fluid sound speed
(4) 20-4-20mA mode
(5) 0-4-20mA mode (can only be used with Version-15 flow meter)
(6)20-0-20mA mode(can only be used with Version-15 flow meter)
(7) 4-20mA corresponding flow velocity
(8)4-20mA corresponding heat flow rate
M56 | 4mA or 0mA output value,
Set the value which corresponds to 4mA or 0mA output current (4mA or 0mA is determined by the setting in M55)
M57 | 20mA output value,
Set the value which corresponds to 20mA output current
M58 | Current loop verification
Check if the current loop is calibrated correctly.
M59 | Display the present output of current loop circuit.
M60 | Setup system date and time. Press ENT for modification. Use the dot key to skip the digits that need no modification.
M61 | Display Version information and Electronic Serial Number (ESN) that is unique for each flow meter.
The users may employ the ESN for instrumentation management
M62 | RS-232/RS485 setup. All the devices connected with flow meter should have matched serial configuration.
The following parameters can be configured: Baud rate (300 to 19200 bps), parity, data bits (always is 8), stop bits (1).
M63 | Select communication protocol.
Factory default is ‘MODBUS ASCII. this is a mode for MODBUS-ASCII, Meter-BUS, Fuji Extended Protocol, Huizhong’s various protocols.
If you are going using MODBUS-RTU you have to select ‘MODBUS_RTU’.
M64 | AI3 value range.
Used to enter temperature/pressure values that are corresponding to 4mA and 20mA input current. The display values have no unit, so that they can present any physical parameter.
M65 | AI4 value range.
Used to enter temperature/pressure values that are corresponding to 4mA and 20mA input current.
M66 | AI5 value range.
Used to enter temperature/pressure values that are corresponding to 4mA and 20mA input current.
M67 | Windows to setup the frequency range (lower and upper limit) for the frequency output function. Valid range is 0Hz-9999Hz. Factory default value is 0-1000 Hz.
For Version-12, Version-13, Version-14 flow meters, you need a hardware module, which shall be plugged to the Serial Expanding Bus, for the frequency output function. Please remember to order the module if you need frequency output function.
For Version-15 flow meter, you need to indicate on your orders that you need the frequency function; Otherwise you will get a flow meter which has no frequency output circuits.
M68 | Window to setup the minimum flow rate value which corresponds to the lower frequency limit of the frequency output.
M69 | Windows to setup the maximum flow Rate value that corresponds to the upper frequency limit of the frequency output.
M70 | LCD display backlight control. The entered value indicates how many seconds the backlight will be on with every key pressing. If the enter value is great than 50000 seconds, It means that the backlight will always keeping on.
M71 | LCD contrast control. The LCD will become darker or brighter when a value is entered.
M72 | Working timer. It can be cleared by pressing ENT key, and then select YES.
M73 | Window to setup the lower limit of flow rate for Alarm#1.
When the flow rate is below the set value, Alarm#1 equals ‘on’
M74 | Window to setup the upper limit of flow rate for Alarm#1.
When the flow rate is above the set value, Alarm#1 equals ‘on’
There are two alarms in the flow meter, and every alarm can be pointed to alarm output devices such as the BUZZER or OCT output or RELAY output. For example, if you want the Alarm#1 is to output by the OCT circuit, you need to set M78 at selection item 6.
M75 | Window to setup the lower limit of flow rate for Alarm#2.
M76 | Window to setup the upper limit of flow rate for Alarm#2.
M77 | Buzzer setup.
If a proper input source is selected, the buzzer will beep when the trigger event occurs. The available trigger sources are:
0. No Signal 1. Poor Signal
2. Not Ready (No*R) 3. Reverse Flow 4. AO Over 100%
5. FO Over 120% 6. Alarm #1 7. Reverse Alarm #2
8. Batch Controller 9. POS Int Pulse 10.NEG Int Pulse
11.NET Int Pulse 12.Energy POS Pulse 13.Energy NEG Pulse
14.Energy NET Pulse 15.MediaVel=>Thresh 16.MediaVelo<Thresh
17.ON/OFF viaRS485 18.Daily Timer (M51) 19.Timed alarm #1
20. Timed alarm #2 21.Batch Total Full 22. Timer by M51
23. Batch 90% Full 24. Key Stroking ON 24.Disable BEEPER
M78 | OCT (Open Collect Transistor Output)/OCT1 setup
By selecting a proper input source, the OCT circuit will close when the trigger event occurs. The available trigger sources are:
0. No Signal 1. Poor Signal
2. Not Ready(No*R) 3. Reverse Flow
4. AO Over 100% 5. FO Over 120%
6. Alarm #1 7. Reverse Alarm #2
8. Batch Controller 9. POS Int Pulse
10.NEG Int Pulse 11.NET Int Pulse
12.Energy POS Pulse 13.Energy NEG Pulse
14.Energy NET Pulse 15.MediaVel=>Thresh
16.MediaVelo<Thresh 17.ON/OFF viaRS485
18. Daily Timer (M51) 19.Timed alarm #1
20. Timed alarm #2 21.Batch Total Full 22.Timer by M51
23.Batch 90% Full 24.Flow Rate Pulse 25.Disable OCT
The OCT circuit does not source voltage at its output. It must be connected with an external power and pull-up resistant for some occasions.
When the OCT circuit is close, it will draw current. The maximum current shall not be over 100mA.
Attention: the maximum voltage applied to OCT can not be over 80 volts.
M79 | Relay or OCT2 setup
By selecting a proper input source, the RELAY will close when the trigger event occurs
The available trigger sources are:
0. No Signal 1. Poor Signal
2. Not Ready(No*R) 3. Reverse Flow 4. AO Over 100%
5. FO Over 120% 6. Alarm #1 7. Reverse Alarm #2
8. Batch Controller 9. POS Int Pulse 10.NEG Int Pulse
11.NET Int Pulse 12.Energy POS Pulse
13.Energy NEG Pulse 14.Energy NET Pulse
15.MediaVel=>Thresh 16.MediaVelo<Thresh
17.ON/OFF viaRS485 18. Timer (M51 Daily)
19.Timed alarm #1 20. Timed alarm #2
21.Batch TotalFull 22.Timer by M51
23.Batch 90% Full 24.Disable RELAY
The RELAY is of SPST(Single pole, single throw) type. It is rated for 110VAC max and have a current rating of 0.5A resistive load.
It highly recommended that a salve relay to be utilized whenever a large resistive load or inductive load is to be controlled.
Note. In order to make the user interface compatible with the former version7, the name RELAY was used other than OCT2, but in fact it is an OCT output.
M80 | Window for selecting the trigger signal for the built-in batch controller. Available trig sources:
0. Key input (press ENT key to start the batch controller)
1. Serial port
2. AI3 rising edge (when AI3 receives 2mA or more current)
3. AI3 falling edge (when AI3 stop receiving 2mA or more current)
4. AI4 rising edge (when AI3 receives 2mA or more current)
5. AI4 falling edge (when AI3 stop receiving 2mA or more current)
6. AI5 rising edge (when AI3 receives 2mA or more current)
7. AI5 falling edge (when AI3 stop receiving 2mA or more current)
8.Timer periodically (define the start time and interval time in M51)
9.Timer daily (define the start time and interval time in M51)
For the input analog current signal, 0 mA indicates “0”, 4mA or more indicates ‘1’.
By selecting item #8, the batch totalizer can be started periodically by the internal timer located at Menu51. When the batch totalizer is full, a signal which indicate the batch is full can be direct to either the OCT or the RELAY terminals to stop the pump or other devices.
By selecting item #9, the batch totalizer could act as totalizer witch runs for only a period of the day so that a alarm signal could be produced if the total flow during that time period is over a certain amount of. For example, if you want a alarm signal which stand for the total flow is over 100 cubic meters during the period of every day from 20:00 to 06:00, setups is like
M51 start time =20:00:00
M51 interval =10:00:00
M51 log times =9999 (means always)
M80 select item #9
M81 input 100 (Unit is defined in M30,M31,M32)
M81 | The built-in batch controller
Set the flow batch value(dose)
The internal output of the batch controller can be directed either to the OCT or the RELAY output circuits.
M81 and M80 should be used together to configure the batch controller.
Note: Because the measuring period is 500mS, the flow for every dos should be keeping at 60 seconds long to get a 1% dose accuracy.
M82 | View the daily, monthly and yearly flow totalizer and thermal energy totalizer value.
The totalizer values and errors for the last 64 days, 32 last 32 months and last 2 years are stored in the RAM memory, To view them, use the ‘ENT’ and ‘UP’ ‘Down’ keys.
M83 | Automatic Amending Function for automatic offline compensation.
Select ‘YES’ to enable this function, select ‘NO’ to disable it.
When the function is enabled, The flow meter will estimate the average flow uncounted (or ‘lost’) during the offline session and add the result to the totalizer.
The estimation of the uncounted flow is made by computing the product of the offline time period and the average flow rate, which is the average of the flow rate before going offline and the one after going on line.
M84 | Set the thermal energy unit:
0. GJ 1. KC 2.KWh 3. BTU
M85 | Select temperature sources
0. from T1,T2 (factory default)
1. from AI3,AI4
M86 | Select the Specific Heat Value.
Factory default is ‘GB’. Under this setting, the flow meter will calculate the enthalpy of water based on the international standard.
If the fluid is other than water, you should select option ‘1. Fixed Specific Heat’, and enter the specific heat value of the fluid.
M87 | Turn on or turn off the Energy totalizer.
M88 | Select thermal energy totalizer multiplying factor.
Factory default is ‘1’.
M89 | 1. Display the temperature difference
2. Window for entering the lowest temperature difference.
M8. | Heat meter is on
1. Inlet
2. Outlet
Select the heat meter installation place.
M90 | Display signal strengths S (one for upstream and one for downstream), and signal quality Q value.
Signal strength is presented by 00.0 to 99.9, the bigger the value, the bigger the signal strength will be, and more reliable readings will be made.
Q value is presented by 00 to 99, the bigger the better. It should at least be great than 50 for normal operations.
M91 | Displays the Time Ratio between the Measured Total Transit Time and the Calculated time. If the pipe parameters are entered correctly and the transducers are properly installed, the ratio value should be in the range of 100±3%. Otherwise the entered parameters and the transducer installation should be checked.
M92 | Displays the estimated fluid sound velocity. If this value has an obvious difference with the actual fluid sound speed, pipe parameters entered and the transducer installation should be checked again.
M93 | Displays total transit time and delta time(transit time difference)
M94 | Displays the Reynolds number and the pipe factor used by the flow rate measurement program. Pipe factor is calculated based on the ratio of the line-average velocity and the cross-section average velocity.
M95 | (1) Display the positive and negative energy totalizers
(2) Upon entering this window, the circular display function will be started automatically. The following windows will be displayed one by one, each window will stay for 8 seconds: M95>>M00>>M01>>M02>>M02>> M03>>M04>>M05>>M06>>M07>>M08>>M90>>M91>>M92>> M93>> M94>>M95. This function allows the user to visit all the important information without any manual action.
To stop this function, simply press a key. Or switch to a window other than M95.
M96 | This is not a window but a command for the thermal printer to advance 5 lines of paper.
M97 | This is not a window but a command to print the pipe parameters.
By default, the produced data will be directed to the internal serial bus (thermal printer). You can also direct those data to the serial communication port.
M98 | This is not a window but a command to print the diagnostic information.
By default, the produced data will be directed to the internal serial bus (thermal printer). You can also direct those data to the serial communication port.
M99 | This is not a window but a command to copy the current display window. By default, the produced data will be directed to the internal serial bus (thermal printer). You can also direct those data to the serial communication port.
By use of the window copying function, you can hardcopy very window displaying manually by switching windows, or you can obtain the window displaying data by communication.
M+0 | Browse the 32 recorded instrument power-on and power-off date and time with the flow rate at the time of power on and off
M+1 | Displays the total working time of the flow meter.
When the backup battery is removed, the total working time will be reset to zero.
M+2 | Displays the last power-off date and time
M+3 | Displays the last power-off flow rate
M+4 | Displays how many times of has been powered on and powered off.
M+5 | A scientific calculator for the convenience of field working.
All the values are in single accuracy.
The calculator can be used while the flow meter is conducting flow measurement.
Water density and PT100 temperature can also be found in this function.
M+6 | Set fluid sound speed threshold
Whenever the estimated sound speed (displayed in M92) exceeds this threshold, an alarms signal will be generated and can transmitted to BUZZER or OCT or RELAY.
This function can used to produce an alarm or output when fluid material changes.
M+7 | Displays total flow for this month(only for the time past)
M+8 | Displays total flow for this year(only for the time past)
M+9 | Display the not-working total time in seconds. The total failure timer will also include the time when power off, if the back-up battery is applied.
M.2 | Entry to solidify the zero point. Password protected.
M.5 | Setup the Q value threshold.
If the present Q is below this threshold, flow rate will be set to 0.
This function is useful when flow meter is installed in noisy environment or on airy pipes.
M.8 | The maximum flow rates for today and this month.
M.9 | Serial port tester with CMM command output for very second.
M-0 | Entry to hardware adjusting windows only for the manufacturer
M-1 | 4-20mA output adjustment
M-2 | 4mA calibration for AI3 input
M-3 | 20mA calibration for AI3 input
M-4 | 4mA calibration for AI4 input
M-5 | 20mA calibration for AI4 input
M-6 | 4mA calibration for AI5 input
M-7 | 20mA calibration for AI5 input
M-8 | Lower Temperature Zero setup for the PT100
M-9 | Higher Temperature Zero setup for the PT100
M-A | Temperature Calibration at 50℃
M-B | Temperature Calibration at 84.5℃

## Sheet2
Flow Transducer | Sensor Body | Transmiter | Model | Measuring Range | Temperature
Clamp on |  | VI-832 / VIR 832M | TS-2 (small) | DN25-100 | -30 ~ 90℃
|  |  | TM-1 (medium) | DN50-700
|  |  | TL-1 (large) | DN300-6000
Insertion |  | VI-832 / VIR 832M | VIR-UF- INSRT-UF | DN50-6000 | -30 ~ 160℃
In-line |  | VI-832 / VIR 832M | VIR-850 | DN15-1000 | -30 ~160℃
Temperature Sensor | Construction | Model | Measuring range | Temperature | Cutoff water
Clamp on |  | CT-1(PT-100) | DN50-6000 | -40 ~ 160℃ | No need
Insertion |  | TCT-1(PT-100) | DN50-6000 | -40 ~ 160℃ | Need
Insertion for Pressurised Pipeline |  | PCT-1(PT-100) | DN50-6000 | -40 ~ 160℃ | No need